## Slide 1: Development of hybrid photomultiplier for Hyper-Kamiokande
Akimichi TAKETA1, Yasuhiro NISHIMURA2
for Hyper-K working group
1ERI, The University of Tokyo
2ICRR, The University of Tokyo

## Slide 2: Purpose: Earth composition measurement
Hyper-K will measure atm-ν oscillation
Oscillation probability depends on electron density, not matter density
By using neutrino oscillation, we can measure the electron density of the medium
  If sterile neutrino does not exist
We have the precise matter density of the earth
  From seismic wave tomography and free oscillation
  They are not direct observation of matter density
Combining matter density and electron density, we can measure the average chemical composition of the deep earth !
  Ratio of atomic number to mass number (Z/A)
Energy and angular resolution is essential for this study
Notes: ポンチ絵

## Slide 3: HPD (Hybrid Photo Detector)
High performance and low cost : better PID and angular resolution
Difficulties to consider viability for practical use in Hyper-K
  Dark noise from AD + Amp., HV around 10kV, low gain, thermal dependence of AD gain, No prior experience using
ｅ
ｅ
ｅ
PMT
HPD
Metal dynode
1-2kV
~8kV
Avalanche Diode (AD)
x ~ 5
x ~400
x ~100
AD gain
Bombardment gain
ΔV ~ -260V
Bombardment gain
Amp
Amp gain
X ~200
Typical values in 8-inch HPD prototype
High voltage around 10kV is required
to collect electrons in small region of AD (~5-10mm) to increase gain at electron-bombardment
 | PMT (20”) | HPD (8”)
HV | 1-2kV | ~8kV
Gain | ~107 | ~104 - 105
C.E. | ~70% | ~97%
Same photo cathode (Q.E.)

## Slide 4: Basic performance of HPD prototype
Charge
[pC]
Pulse height distribution
Amplitude [mV]
1 p.e.
2 p.e.
3 p.e.
Signal timing of 1 - 4 photons
Time Transition Spread (TTS) [ns]
photo electrons (p.e.)
Linearity
σ[%]
σ [%]
up to 3 p.e.
Intrinsic resolution of only HPD
620 ps σ /p.e.
(With 70m cable)
All results are preliminary by HPD prototype.

## Slide 5: Overview of HPD study
0.2% Gd water in 200 ton water tank
20” PMTs x 240
Pre-treatment system
200t tank
7 m
Replaced to 8 of 8-inch HPDs
Development of HPD with Hamamatsu photonics
Proof test in EGADS tank (2013 - )
Evaluating Gadolinium’s Action on Detector Systems
Calibration (March 2013)
Precise performance evaluation
Gain adjustment of 10 HPDs
Safety test in water
2 HPDs are evaluated in detail.
→ Finally test 20” HPD in 200t tank based on 8”HPD R&D
8” HPD
(8” PMT)

## Slide 6: 8-inch HPD
Preamp (under development)
Control
Coaxial
LV PS
AD out / Amp in
Amp out
HPD
cable in water
項目 | 値
感度波長範囲 | 300 ～ 650 nm
光電面 | バイアルカリ
面板材質 | 硼珪酸
２次電子増倍方式 | 半導体電子打ち込み式
ターゲット半導体 | 5mmφアバランシェダイオード
20 ns
1 p.e.
10 mV
rise / fall 8 / 16 ns
Pulse width : 18ns
HPD(8-inch)
PMT(20-inch)
Signal
Range of sensitive wavelengths |  | 300 - 650 nm
Typical voltage | Photocathode | 8 kV
 | AD bias | 260 V
Gain |  | 4ｘ104
Single-photon resolution（σ） |  | 20 %
Transition time spread（σ） |  | 620 ps
8-inch photocathode
High voltage module (2ch 10kV/500V Max)
LV
10V
Signal
5mm φ
avalanche diode (AD)
Cable
30cm
20cm
Hamamatsu Photonics
filter
filter
（+8kV）
（+8kV - ΔV）
AD
Inst. amp
100 ns
I-V amp.
10ch control power supply
Signal from amplifier
No HV line in water!

## Slide 7: Hit threshold and efficiency
Better trigger efficiency is expected thanks to narrow 1 p.e. peak at 8” HPD.
  Even if threshold is higher than 0.25 p.e.
Calibration setup and gain, threshold optimization in Feb 2013.
1 photon events
1 p.e. signal
pedestal
Fitting function
and each component
20” PMT
1 p.e.
← pedestal
Signal efficiency[%]
= Sig.(>threshold) / Sig.(All)
Threshold [mV]
Purity[%]
= {Sig. /(Sig. + Ped.)} (>threshold)
Trigger efficiency [%]
0.25 p.e. threshold is set for 1 hit at Super-K (20” PMT)
0.25p.e.→
1 p.e.→
0.25p.e.
8” HPD
Ex trigger

## Slide 8: HPD performance
Dynamic range is wide by HPD itself, but limited by amplifier currently mounted on 8”HPD.
  3-60mV/p.e. from 8” HPD
  Linearity up to ~600mV
    < ~200p.e.
    Sufficient if amplitude gain is same as SK-PMT.
  Amplifier will be developed continuously.
P/V ratio and TTS σ is limited by Amplifier
  Intrinsic noise
  Trans impedance problem
~600mV
Pulse height - charge

## Slide 9: Potential problem of 20-inch HPD
filter
filter
（+8kV）
（+8kV - ΔV）
AD
Inst. amp
I-V amp.
Larger Parasitic capacitance of AD leads:
  worse frequency response
  depends on I-V gain (= amp impedance)
If we reduce impedance
  We will loose gain = P/V ratio, TTS σ
  Inst-amp cannot recover this loss because of current noise density. (It’s floating measurement)

## Slide 10: Possible solution
Low impedance of 1st full differential amp
  Parasitic capacitance is negligible
  Better frequency response
Multiplex return
  Reduce trans-impedance amplifier noise
Single power supply (+5V)

## Slide 11: Current status (concludion)
Production of ten 8-inch HPDs has finished. for proof test in 200t tank.
  Long time running, durability of HV switching and safety operation in water were checked.
  Quality test was finished
  Proof test will start in this summer
1st 20-inch HPD will be produced in this year
2013
2015
2016
2014
Jan
Feb
Mar
Apr
8”HPD R&D
Ten 8” HPD production
Proof test ( tank )
Calibration
Durability test
Test
20”HPD R&D
Performance evaluation
200 t
Installation
2
8
Relative to 20”PMT,
trigger performance,
durability, stability,..
Uniformity, band width, rate tolerance, thermal・B-field dependence, dark rate..
1
2
3