## Slide 1
Scanning Microscope for muon radiography with nuclear emulsion
Cristiano Bozza1, Lucia Consiglio2, Nicola D'Ambrosio3, Giovanni De Lellis4, Chiara De Sio5, Natalia Di Marco3, Umut Kose6, Eduardo Medinaceli7, Seigo Miyamoto8, Ryuichi Nishiyama8, Fabio Pupilli3, Simona Maria Stellacci1, Chiara Sirignano7, Paolo Strolin4, Hiroyuki Tanaka8, Valeri Tioukov2
University of Salerno and INFN1, INFN Napoli2, INFN / LNGS3, University of Napoli and INFN4, University of Salerno5, INFN Padova6, University of Padova and INFN7, University of Tokyo8

## Slide 2
Scanning Microscope
Nuclear emulsion detectors for muon radiography
Detectors are made of stacked emulsion films
Emulsion has no time resolution, no trigger: all tracks are recorded
m
m
e+e-
e+e-
e+e-
Emulsion films record hard tracks as well as soft tracks
3D information available for each track: momentum discrimination and/or particle id. possible!
for muon radiography with nuclear emulsion

## Slide 3
Scanning Microscope
Nuclear emulsion images
1 μm
Charged particles ionize Ag atoms (stochastic process), producing the latent image
AgBr gel
Metallic Ag grows in filaments during development
With green-white light the average l is 600 nm: the filaments cannot be resolved because of diffraction
“Grains” = clusters of filaments
for muon radiography with nuclear emulsion

## Slide 4
Scanning Microscope
for muon radiography with nuclear emulsion
Looking at emulsion films: basic optical setup
Emulsion film
Lamp (optionally w/ filters)
White, green or blue
Plastic base
Condenser lens
Illuminated spot
Objective lens (or lens system)
CMOS sensor

## Slide 5
Scanning Microscope
Nuclear emulsion images
Imaging by objective + camera: the spatial density of metallic Ag is folded with the PSF (point-spread function), characterizing the optical setup
Y(x,y,z)
Focal plane
Out of focus
Out of focus
Typical grain size after development: 0.2÷1 μm
(0.5 μm in the case shown in this talk)
Grains in emulsion image: high-energy tracks, electrons, fog (randomly developed grains, not touched by any ionizing particle)
50 μm
Depth of field: ~3 μm
for muon radiography with nuclear emulsion

## Slide 6
Scanning Microscope
Nuclear emulsion images
Grain images are not uniform, and depend more on the neighborhoods than on the features of grains themselves
Finite depth of field: grains out of focus can be seen
Shadow effect: grains stacked one on top of the other on the focal axis look darker (bigger)
Highly ionizing particles: grains may be so close they cannot be resolved
A “hole”: no doubt the charged particle passed there, but it just did not ionize!
for muon radiography with nuclear emulsion

## Slide 7
Scanning Microscope
Nuclear emulsion images
for muon radiography with nuclear emulsion
3D tomography: change focal plane
Alignment residuals of track grains: 50 nm in optical microscopy!
Good precision helps rejecting random alignments and thus increasing signal/background ratio

## Slide 8
Scanning Microscope
The European Scanning System (ESS)
for muon radiography with nuclear emulsion
Developed for OPERA, used in all European laboratories
Also installed at Tokyo ERI
Scanning speed: 20 cm2/h/side
CMOS camera 1280×1024 pixel 256 gray levels 376 frames/sec (Mikrotron MC1310)
XY stage (Micos) 0.1 μm nominal precision
Emulsion Plate
Z stage (Micos) 0.05 μm nominal precision
Illumination system, objective (Oil 50× NA 0.85) and optical tube (Nikon)

## Slide 9
Scanning Microscope
The ESS: working principles
for muon radiography with nuclear emulsion
280×365μm2
Tomographic sequences Z axis moving, 2D images spanning emulsion thickness
Move XYZ to next view
Process/save data
Next field of view, Z at top, new cycle
DAQ cycle (185 ms)
Camera
2D Images (peak 452 MB/s, avg. 97 MB/s)
Vision Processor (Matrox Odyssey)
Binarized 2D Images
Host PC (Dual Pentium Workstation) Running WinXP
Grains
XYZ Motion Commands
Motion Controller (National Instruments FlexMotion)
Motors (VEXTA Nanostep)
Power
Functional blocks

## Slide 10
Scanning Microscope
The ESS:
Image processing
(SySal2000)
for muon radiography with nuclear emulsion
15 images 10 grains signal/image, 3000 grains background+ noise, shadows, scratches, spots
2D FIR Filter+ Equalization+
Threshold
Grain recognition (Host PC, multithreaded Assembler code)
300 ÷ 3000 microtracks / view
3D microtrack reconstruction (Host PC or tracking servers, multithreaded C++ code)

## Slide 11
Scanning Microscope
for muon radiography with nuclear emulsion
Tracking: recognition of aligned sequences of grains in 2D images (microtracks)
Highly optimized algorithm to deal with big combinatorial complexity
Parallel processing: can use up to 8 processors/cores per machine
The ESS: Tracking (SySal2000)

## Slide 12
Scanning Microscope
The ESS: Tracking (SySal2000)
for muon radiography with nuclear emulsion
ScanGrid: use powerful machines dedicated to on-line tracking/computing and simplify the architecture of data-taking
Grains
tracks
Other info
(setup, monitoring)
Chart: Time(s)
Installation in Salerno: 70 tracking cores shared by 4 microscopes
Installation at LNGS: 80 tracking cores shared by 10 microscopes
Automatic load balancing (different quality of emulsion requires different processing power)

## Slide 13
Scanning Microscope
for muon radiography with nuclear emulsion
Tests on 8 GeV/c pion beams
The ESS: current performances
Sy = 0
Sy = -0.180
Base-track
Microtrack
Notice: efficiency depends on emulsion quality!!!

## Slide 14
Scanning Microscope
for muon radiography with nuclear emulsion
Precision of film-to-film track connection
The ESS: current performances
Sx = 0.025 Sy = 0
Sx = 0.600 Sy = -0.180

## Slide 15
Scanning Microscope
for muon radiography with nuclear emulsion
Precision of film-to-film track connection
The ESS: current performances
Sx = 0.025 Sy = 0
Sx = 0.600 Sy = -0.180

## Slide 16
Scanning Microscope
The Quick Scanning System (QSS): evolution of the ESS
for muon radiography with nuclear emulsion
Increase sensor size (area scanned at each tomographic sweep)
Increase grabbing speed
Increase processing speed
Reduce dead time due to motion
Increase scanning speed: enable using larger areas  higher statistics
improve signal/background ratio
improve sensitivity to flux variations
improve sensitivity to density variations
Keep data quality high:
low number of fake tracks  effectively discriminate electrons from muons
good precision  effectively discriminate electrons from muons
high efficiency  increase data rate/unit area (with triplet or quadruplet stacks, microtracking efficiency suppresses statistics with at least the 6th power!)

## Slide 17
Scanning Microscope
The QSS
for muon radiography with nuclear emulsion
Same mechanics, new hardware
4 Mpixel camera, 400 fps
Double CL frame grabber (Matrox Radient)
New optics (20×)
Image processing and tracking by GPU
New motion control unit
Pro-Dex MAXk

## Slide 18
Scanning Microscope
The Quick Scanning System (QSS): evolution of the ESS
for muon radiography with nuclear emulsion
Z motion, 2D images read and processed on-the-fly
(vision processor in host PC)
“Stop’n’go”
2D images read and processed on-the-fly (GPU in host PC)
Dead time due to motion is reduced
“Continuous motion” X axis travels at constant speed

## Slide 19
Scanning Microscope
The Quick Scanning System (QSS): evolution of the ESS
for muon radiography with nuclear emulsion
Steps for track reconstruction
Dark pixel clustering
Image-to-image alignment (same view)
Chains of clusters
(1 chain = 1 or more grains)
View-to-view alignment by chain pattern matching
Microtrack recognition
Base-track linking (uses standard SySal.NET linking)
GPU
GPU
GPU
GPU
CPU
GPU

## Slide 20
Scanning Microscope
The Quick Scanning System (QSS): evolution of the ESS
for muon radiography with nuclear emulsion
Image-to-image alignment
1 ms stage sampling loop on XYZ (does not require piezodrive)
Image distortion corrections
XY curvature
Z curvature
Z axis slant (X and Y)
XY Trapezium
Magnification vs. Z

## Slide 21
Scanning Microscope
The Quick Scanning System (QSS): evolution of the ESS
for muon radiography with nuclear emulsion
Image-to-image alignment results
mm
mm
XY precision: 0.12 mm

## Slide 22
Scanning Microscope
The Quick Scanning System (QSS): evolution of the ESS
for muon radiography with nuclear emulsion
3D track recognition:
Microtracks: sequences of aligned grains (230 nm tolerance on X/Y, 2 μm on Z) recognized in the whole scanning volume
GPU-based tracking server (2× NVidia GTX690) – 3072 CUDA cores/microscope
View-to-view mapping: discard duplicated grains in overlap region and correct misalignments due to stage motion

## Slide 23
Scanning Microscope
The Quick Scanning System (QSS): evolution of the ESS
for muon radiography with nuclear emulsion
View-to-view 3D pattern matching
Recovers transverse vibrations and XYZ sampling errors (allows microtracking across views)
Merges chain duplicates in overlap volume (prevents excess microtracks)
XY precision: 0.2 mm
Z precision: 2.6 mm
mm
mm
mm

## Slide 24
Scanning Microscope
The Quick Scanning System (QSS): evolution of the ESS
for muon radiography with nuclear emulsion
Tracking on test films exposed to large angle beams
Preliminary!

## Slide 25
Scanning Microscope
The Quick Scanning System (QSS): evolution of the ESS
for muon radiography with nuclear emulsion
Tracking on test films exposed to large angle beams
Preliminary!

## Slide 26
Scanning Microscope
The Quick Scanning System (QSS): evolution of the ESS
for muon radiography with nuclear emulsion
Depending on conditions, it is possible to speed up the system with minimal changes (tuning quantities in the parameter form)
The scanning speed is 41 cm2/h with 31 layers, 58% view travel, 200 fps operation
This should ensure maximum signal/noise ratio
3121 layers (glued emulsions, low fog)  64 cm2/h (tested)
+ 58%75% view travel (8 core CPU)  85 cm2/h (technically tested)
+ 200400 fps  120 cm2/h (estimated)
For future applications, having low fog emulsion will in general improve the performances (fewer layers needed for high efficiency, low combinatorial background)
Cost of the upgrade from ESS to QSS: about 20 k€!

## Slide 27
Scanning Microscope
Conclusions
for muon radiography with nuclear emulsion
The ESS is an established technology that provides the performances needed for muon radiography
Film-to-film connection at micrometric level (slope accuracy of the order of 10 mrad is easily achieved)
The QSS is the upgraded project that is approaching its first stage (2× speed) with cheap hardware upgrades
Outlook: 6× speed increasing just the number of GPU’s (4 × 600 €)
Stay tuned for first applications of the QSS in muon radiography!
The ESS has been used for Unzen and Stromboli data readout